## BEMSdar
MINISTRY OF HEALTH
Deduction Mapping Report Biomedical Engineering Maintenance Services (BEMS)
|  |  | HOSPITAL NAME : | Hospital Orang Asli Gombak | STATE: | W.P. Kuala Lumpur
|  |  | MONTH/YEAR : | 2026-05-01 | COMPANY: | Ambang Wira Sdn Bhd
MONTHLY SERVICE FEE (RM) : |  |  |  | 17762.21
BEMS MONTHLY FEE (RM) @ 97.5% : |  |  |  | 17318.1547
FMS MONTHLY FEE (RM) @ 2.5% : |  |  |  | 444.0553
No | Key Deduction Indicators |  |  |  | Frequency |  | Deduction Value (RM) | Deduction % | Performance Level (%)
B.1 | All service requests shall be responded by the Company’s trained technical personal within the timeframe as stipulated in the Agreement. |  |  |  | Per event |  | 0 | 0 | 100
B.2 | All unscheduled maintenance works including relevant maintenance calibration, safety tests and functional checks shall be completed within seven (7) working days. |  |  |  | Per day |  | 0 | 0 | 100
B.3 | Scheduled maintenance shall be carried out as per schedule and according to the procedure as specified in the Agreement. Scheduled maintenance shall include complete safety and performance tests. |  |  |  | Per event |  | 0 | 0 | 100
B.4 | All biomedical equipment shall be functioning and maintained to ensure uptime target as specified in the TRKPI are met. |  |  |  | Per event |  | 0 | 0 | 100
B.5 | T&C and Warranty Management of newly introduced biomedical shall be managed and performed according to the Agreement and related acts, standards and regulations. |  |  |  | Per event |  | 0 | 0 | 100
Total | [6 empty] | 0 | 0 | 100
|  | Prepared by: |  |  | Verified by:
|  | NUR ALIAH BINTI HUSSIN |  |  | JULIWATI BINTI MATSAH
|  | Head of BEMS |  |  | Facility Manager
|  | Ambang Wira Sdn Bhd- Hospital Orang Asli Gombak |  |  | Ambang Wira Sdn Bhd- Hospital Orang Asli Gombak
|  | Date: |  |  | Date:
Note on E9: USER: fee strands: RM1,163,338.87

## B1-RESPONSE & B2-REPAIR
MINISTRY OF HEALTH
Deduction Mapping Report Biomedical Engineering Maintenance Services (BEMS)
INDICATOR B.1 RESPONSE TIME & B.2 REPAIR TIME
HOSPITAL NAME: |  | HOSPITAL ORANG ASLI GOMBAK | [5 empty] | FIRST DATE OF THE MONTH : |  |  | 1/5/2026 |  |  |  | (For Calculation)
STATE: |  | WILAYAH PERSEKUTUAN KUALA LUMPUR | [5 empty] | LAST DATE OF THE MONTH : |  |  | 31/5/2026 |  |  |  | (For Calculation)
COMPANY: |  | AMBANG WIRA SDN BHD | [5 empty] | BEMS MONTHLY FEE (RM) @ 97.5% : |  |  | 17318.1547
MONTH/YEAR: |  | 2026-05-01 | [5 empty] | FMS MONTHLY FEE (RM) @ 2.5% : |  |  | 444.0553
Notes:
1) Deduction Ringgit Equivalent is based on Appendix 11 of the Agreement, Section B3 Key Deduction Indicators for BEMS
2) Deduction Value for Indicator B.1 is Deduction Ringgit Equivalent (per event)
3) Total Repair Time includes weekend & public holiday
4) Demerit Point for Indicator B.2 is calculated based on completion date against request date, less by 7 working days
5) Deduction Value for Indicator B.2 is Demerit Point multiply by Deduction Ringgit Equivalent
No | Asset No | Asset Description | Asset Purchase Price (RM) | Work Order No. | User Location | Work Order/ Complaint Details | Response Category (Emergency/ Normal) | Total Response Time (minutes) | Work Request Date | Work Completion Date | Deduction Ringgit Equivalent (RM) | Indicator B.1 |  |  | Indicator B.2
[12 empty] | Validate Status (Y/N) | Deduction Value (RM) | Remarks/ Exemption | Validate Status (Y/N) | Total Repair Time (day) | Last date of 7 working days (dd/mm/yy) | Demerit Point | Deduction Value (RM) | Remarks/ Exemption
Current Month (Work Order Completed)
1 | HOAG 10309 | ECG (EKG) / EPG PROJECT - ECG MACHINE | 5950 | HOAGWIBE26003347 | BILIK RAWATAN (WAD EMPAT) | PENYEPIT ECG MACHINE PATAH | N | 5.0049 | 2026-05-09 | 2026-05-11 | 20 | N | 0 |  | N | 3 | 2026-05-18 | 0 | 0 | Breakdown
2 | HOAG 10019 | PATIENT MONITORING EQUIP (MINDRAY) | 5795 | HOAGWIBE26003595 | RUANG RED ZONE | BP MONITOR PROBLEM | N | 5.0139 | 2026-05-13 | 2026-05-13 | 0 | N | 0 |  | N | 1 | 2026-05-21 | 0 | 0 | Breakdown
3 | HOAG 10420 | DEFIBRILLATOR | 75015.73 | HOAGWIBE26003617 | BANGUNAN KECEMASAN & KEMALANGAN | MESIN DEFIBRILLATOR TIDAK BOLEH CALIBRATION | N | 5.0215 | 2026-05-16 | 2026-05-16 | 100 | N | 0 |  | N | 1 | 2026-05-25 | 0 | 0 | Breakdown
4 | HOAG 10063 | SPOT VITAL SIGN / TRANSPORT VITAL SIGN MONITOR | 8000 | HOAGWIBE26003784 | BILIK RAWATAN 1 (UNIT OPD) | BP SET MONITOR TIDAK NAIK | N | 5.0069 | 2026-05-20 | 2026-05-20 | 20 | N | 0 |  | N | 1 | 2026-05-29 | 0 | 0 | Breakdown
5 | HOAG 10132 | ADEC OFFICER UNIT DENTAL CHAIR AND STOOL | 49500 | HOAGWIBE26003881 | BILIK SURGERI 3 (UNIT DENTAL) | HANDPIECE & SCALER UNTUK DENTAL CHAIR TIDAK KELUAR AIR | N | 5.0167 | 2026-05-25 | 2026-05-25 | 100 | N | 0 |  | N | 1 | 2026-06-05 | 0 | 0 | Breakdown
6 | HOAG 10420 | DEFIBRILLATOR | 75015.73 | HOAGWIBE26003883 | BANGUNAN KECEMASAN & KEMALANGAN | CM TUKAR BATTERY DEFIBRILLATOR | N | 5.0194 | 2026-05-25 | 2026-05-26 | 100 | N | 0 |  | N | 2 | 2026-06-05 | 0 | 0 | Corrective
7 | HOAG 10369 | PATIENT MONITORING EQUIPMENT | 8500 | HOAGWIBE26003886 | BILIK RAWATAN (WAD TIGA) | CUFF BP MESIN MINDRAY KOYAK | N | 5.0083 | 2026-05-26 | 2026-05-26 | 20 | N | 0 |  | N | 1 | 2026-06-08 | 0 | 0 | Breakdown
8 | HOAG 10289 | PATIENT MONITORING EQUIPMENT | 3000 | HOAGWIBE26003896 | WAD TIGA | WAYAR UNTUK PEMERIKSA OKSIGEN DALAM DARAH TIDAK BERFUNGSI | N | 5.0201 | 2026-05-29 | 2026-05-29 | 10 | N | 0 |  | N | 1 | 2026-06-10 | 0 | 0 | Breakdown
9 | HOAG 10202 | SPOT VITAL SIGN MONITOR SPO2 (ANGELUS) | 13650 | HOAGWIBE26003897 | BILIK RAWATAN (WAD TIGA) | BP CUFF KOYAK | N | 5.009 | 2026-05-29 | 2026-05-29 | 30 | N | 0 |  | N | 1 | 2026-06-10 | 0 | 0 | Breakdown
Current Month (Work Order Open)
Previous Month (Work Order Closed)
1 | HOAG 10108 | ADEC CASCADE BASIC 2140 AND STOOL | 49950 | HOAGWIBE26002911 | UNIT DENTAL | SLOW SPEED BUR TIDAK PUSING | N | 5.0125 | 2026-04-23 | 2026-05-04 | 100 | N | 0 |  | N | 12 | 2026-05-04 | 0 | 0 | Breakdown
Previous Month (Work Order Open)
TOTAL DEDUCTION FOR IND. B.1: | [12 empty] | 0 | TOTAL DEDUCTION FOR IND. B.2: |  |  |  |  | 0
[9 empty] | 2019-09-10 08:50 | 2019-09-30 23:59 | 495:09 | 94.35%
Formulas:
Last date of 7 working days (dd/mm/yy) = WORKDAY([Work Request Date],6,tarikhcuti)

## B3-SCHEDULED
MINISTRY OF HEALTH
Deduction Mapping Report Biomedical Engineering Maintenance Services (BEMS)
INDICATOR B.3 SCHEDULED MAINTENANCE
HOSPITAL NAME : |  | HOSPITAL ORANG ASLI GOMBAK |  |  |  |  | MONTH/YEAR : |  | 2026-05-01
| STATE : | WILAYAH PERSEKUTUAN KUALA LUMPUR |  |  | BEMS MONTHLY FEE (RM) @ 97.5% : |  |  |  | 17318.1547
| COMPANY : | AMBANG WIRA SDN BHD |  |  | FMS MONTHLY FEE (RM) @ 2.5% : |  |  |  | 444.0553
Notes:
1) Deduction Ringgit Equivalent is based on Appendix 11 of the Agreement, Section B3 Key Deduction Indicators for BEMS
2) Deduction Value for Indicator B.3 is Deduction Ringgit Equivalent (per event)
3) Rescheduled Date is the final rescheduled date if there are multiple rescheduling
4) Display only asset affected by deduction
No | Asset No | Asset Description | Work Order No. | User Location | Scheduled Start Date | Scheduled End Date | Rescheduled Date | Maintenance Date | Validate Status (Y/N) | Deduction Ringgit Equivalent (RM) | Indicator B.3
[11 empty] | Deduction Value (RM) | Remarks/ Exemption
1 | KKM/JKN/HOAG/10-01-01/I/14/257 | THERMOMETER, ELECTRONIC | HOAGWOBE26003227 | BILIK UBAT | 2026-05-08 | 2026-05-09 |  | 2026-05-08 | N |  |  | PPM DONE
2 | HOAG 10400 | BODY COMPOSITION MONITOR WEIGHING | HOAGWOBE26003228 | UNIT KUALITI (UNIT SPKAKOA) | 2026-05-08 | 2026-05-09 |  | 2026-05-08 | N |  |  | PPM DONE
3 | KKM/JKN/HOAG/10-01-01/I/14/149 | SCALES, PATIENT, FLOOR | HOAGWOBE26003229 | BILIK PEGAWAI SAINS PEMAKANAN (UNIT SPKAKOA) | 2026-05-14 | 2026-05-15 |  | 2026-05-14 | N |  |  | PPM DONE
4 | KKM/JKN/HOAG/10-01-01/R/17/30 | SCALES, PATIENT, FLOOR | HOAGWOBE26003230 | BILIK PEGAWAI SAINS PEMAKANAN (UNIT SPKAKOA) | 2026-05-14 | 2026-05-15 |  | 2026-05-14 | N |  |  | PPM DONE
5 | KKM/JKN/HOAG/10-01-01/R/18/24 | SPHYGMOMANOMETERS, ELECTRONIC, AUTOMATIC | HOAGWOBE26003231 | BILIK PENYELIA HOSPITAL & PENYELIA JURURAWAT (UNIT SPKAKOA) | 2026-05-15 | 2026-05-16 |  | 2026-05-15 | N |  |  | PPM DONE
6 | KKM/JKN/HOAG/10-01-01/H/13/51 | SPHYGMOMANOMETERS, ELECTRONIC, AUTOMATIC | HOAGWOBE26003232 | PANTRI (UNIT SPKAKOA) | 2026-05-15 | 2026-05-16 |  | 2026-05-15 | N |  |  | PPM DONE
7 | KKM/JKN/HOAG/10-01-01/R/17/28 | SPHYGMOMANOMETERS, ELECTRONIC, AUTOMATIC | HOAGWOBE26003233 | PANTRI (UNIT SPKAKOA) | 2026-05-15 | 2026-05-16 |  | 2026-05-15 | N |  |  | PPM DONE
8 | KKM/JKN/HOAG/10-01-01/R/17/27 | SPHYGMOMANOMETERS, ELECTRONIC, AUTOMATIC | HOAGWOBE26003234 | PANTRI (UNIT SPKAKOA) | 2026-05-15 | 2026-05-16 |  | 2026-05-15 | N |  |  | PPM DONE
9 | KKM/JKN/HOAG/10-01-01/I/13/279 | SPHYGMOMANOMETERS, ELECTRONIC, AUTOMATIC | HOAGWOBE26003235 | PANTRI (UNIT SPKAKOA) | 2026-05-15 | 2026-05-16 |  | 2026-05-15 | N |  |  | PPM DONE
10 | KKM/JKN/HOAG/10-01-01/I/13/280 | SPHYGMOMANOMETERS, ELECTRONIC, AUTOMATIC | HOAGWOBE26003236 | PANTRI (UNIT SPKAKOA) | 2026-05-15 | 2026-05-16 |  | 2026-05-15 | N |  |  | PPM DONE
11 | KKM/JKN/HOAG/10-01-01/R/17/25 | ANALYZERS, POINT-OF-CARE, WHOLE BLOOD, GLUCOSE | HOAGWOBE26003237 | PANTRI (UNIT SPKAKOA) | 2026-05-15 | 2026-05-16 |  | 2026-05-15 | N |  |  | PPM DONE
12 | KKM/JKN/HOAG/10-01-01/H/17/6 | VENTILATORS, TRANSPORT | HOAGWOBE26003238 | RUANG RED ZONE | 2026-05-18 | 2026-05-19 |  | 2026-05-18 | N |  |  | PPM DONE
13 | KKM/JKN/HOAG/10-01-01/H/13/66 | SCANNING SYSTEMS, ULTRASONIC, ABDOMINAL | HOAGWOBE26003239 | BILIK RAWATAN | 2026-05-19 | 2026-05-20 |  | 2026-05-19 | N |  |  | PPM DONE
14 | KKM/JKN/HOAG/10-01-01/H/14/75 | ASPIRATORS | HOAGWOBE26003240 | BILIK RAWATAN | 2026-05-22 | 2026-05-23 |  | 2026-05-22 | N |  |  | PPM DONE
15 | KKM/JKN/HOAG/10-01-01/H/09/21 | OPHTHALMOSCOPES | HOAGWOBE26003241 | KLINIK OFTAMOLOGI (UNIT OPD) | 2026-05-27 | 2026-05-28 |  | 2026-05-26 | N |  |  | PPM DONE
16 | KKM/JKN/HOAG/10-01-01/H/05/17 | LARYNGOSCOPES, RIGID | HOAGWOBE26003242 | KLINIK OFTAMOLOGI (UNIT OPD) | 2026-05-27 | 2026-05-28 |  | 2026-05-26 | N |  |  | PPM DONE
17 | HOAG 10434 | PHOTOTHERAPY MACHINE | HOAGWOBE26003268 | WAD 2 | 2026-05-06 | 2026-05-07 |  | 2026-05-06 | N |  |  | PPM DONE
Current Month (Open)
Previous Month (Closed)
Previous Month (Open)
TOTAL DEDUCTION FOR IND. B.3: | [10 empty] | 0

## B4-UPTIME
MINISTRY OF HEALTH
Deduction Mapping Report Biomedical Engineering Maintenance Services (BEMS)
INDICATOR B.4 UPTIME GUARANTEE
| HOSPITAL NAME : | HOSPITAL ORANG ASLI GOMBAK | [10 empty] | MONTH/YEAR : | [6 empty] | 2026-05-01
| STATE : | WILAYAH PERSEKUTUAN KUALA LUMPUR | [10 empty] | BEMS MONTHLY FEE (RM)@ 97.5% : | [6 empty] | 17318.1547
| COMPANY : | AMBANG WIRA SDN BHD | [10 empty] | FMS MONTHLY FEE (RM) @ 2.5% : | [6 empty] | 444.0553
Notes:
1) Deduction Ringgit Equivalent is based on Appendix 11 of the Agreement, Section B3 Key Deduction Indicators for BEMS
2) Deduction Value for 1st cut off is Deduction Ringgit Equivalent (per event)
3) Deduction Value for 2nd cut off is Deduction Ringgit Equivalent (per event) multiply by 2
4) Display only asset affected by deduction
No. | Asset No | Type Code | Type Description | Asset Purchase cost | Asset Age | TRKPI % | Uptime % | [11 empty] | Deduction Ringgit Equivalent (RM) | Previous Deduction (Month) |  | Deduction for the month (Y/N) |  | Deduction Value (RM) |  | Total Deduction (RM) | Remarks
[7 empty] | Jan | Feb | Mar | Apr | May | Jun | Jul | Aug | Sep | Oct | Nov | Dec |  | 1st | 2nd | 1st | 2nd | 1st | 2nd
1 | HOAG 10043 | 17-882 | DEFIBRILLTOR ZOLL | 65000 | 5-10 YRS | 95% | Uptime % Jan: 99.98%; Uptime % Mar: 99.38% | 140 |  |  | N | N | 0 | 0 | 0
2 | HOAG 10103 | 17-740 | HEMATOLOGY ANALYSER, | 189000 | >10 | 85% | Uptime % Jan: 99.98% | 380 |  |  | N | N | 0 | 0 | 0
3 | HOAG 10242 | 12-276 | ANGLE POISE LAMP | 150 | >10 | 85% | Uptime % Jan: 99.99% | 10 |  |  | N | N | 0 | 0 | 0
4 | HOAG 10001 | 11-407 | ELECTROCARDIOGRAPHS | 11990 | 5-10 YRS | 92% | Uptime % Jan: 99.99% | 30 |  |  | N | N | 0 | 0 | 0
5 | HOAG 10414 | 10-551 | FUNDUS CAMERA | 97000 | < 5YRS | 96% | Uptime % Jan: 99.9% | 200 |  |  | N | N | 0 | 0 | 0
6 | HOAG 10200 | 12-636 | PATIENT MONITORING EQUIP. PHILLIP GOLDWAY | 24999 | >10 | 88% | Uptime % Jan: 99.99%; Uptime % Apr: 99.98% | 50 |  |  | N | N | 0 | 0 | 0
7 | HOAG 10035 | 16-495 | INFUSION PUMP | 3850 | >10 | 88% | Uptime % Jan: 99.99% | 10 |  |  | N | N | 0 | 0 | 0
8 | HOAG 10034 | 16-495 | INFUSION PUMP | 3850 | >10 | 88% | Uptime % Jan: 99.99% | 10 |  |  | N | N | 0 | 0 | 0
9 | HOAG 10108 | 11-165 | DENTAL DELIVERY UNIT | 49950 | >10 | 85% | Uptime % Jan: 99.09%; Uptime % Feb: 97.79%; Uptime % Mar: 97.64%; Uptime % Apr: 97.21%; Uptime % May: 97.16% | 100 |  |  | N | N | 0 | 0 | 0
10 | HOAG 10373 | 13-217 | INFUSION PUMP | 1400 | < 5YRS | 99% | Uptime % Jan: 99.96% | 10 |  |  | N | N | 0 | 0 | 0
11 | HOAG 10137 | 15-612 | WATER DISTILLER | 7000 | >10 | 85% | Uptime % Feb: 99.71% | 20 |  |  | N | N | 0 | 0 | 0
12 | HOAG 10345 | 12-815 | OPTHALMOSCOPE | 1350 | >10 | 85% | Uptime % Feb: 99.47% | 10 |  |  | N | N | 0 | 0 | 0
13 | HOAG 10217 | 11-134 | DEFIBRILLATOR | 12490 | >10 | 88% | Uptime % Feb: 99.99% | 30 |  |  | N | N | 0 | 0 | 0
14 | HOAG 10277 | 11-134 | DEFIBRILLATOR | 12490 | >10 | 88% | Uptime % Feb: 99.99% | 30 |  |  | N | N | 0 | 0 | 0
15 | HOAG 10132 | 11-165 | DENTAL DELIVERY UNIT | 49500 | >10 | 85% | Uptime % Feb: 99.71%; Uptime % Mar: 98.22%; Uptime % Apr: 98.12%; Uptime % May: 98.08% | 100 |  |  | N | N | 0 | 0 | 0
16 | HOAG 10125 | 20-000 | PORTABLE CUTTING UNIT | 29500 | >10 | 85% | Uptime % Mar: 99.47% | 60 |  |  | N | N | 0 | 0 | 0
17 | HOAG 10187 | 16-241 | ULTRASOUND SCANNER | 40000 | 5-10 YRS | 92% | Uptime % Mar: 99.87% | 90 |  |  | N | N | 0 | 0 | 0
18 | HOAG 10420 | 11-134 | DEFIBRILLATOR | 75015.73 | < 5YRS | 99% | Uptime % Mar: 99.99%; Uptime % Apr: 99.98%; Uptime % May: 99.94% | 160 |  |  | N | N | 0 | 0 | 0
19 | HOAG 10297 | 10-208 | SUCTION PUMP (ASPIRATOR) | 4490 | >10 | 88% | Uptime % Apr: 99.98% | 10 |  |  | N | N | 0 | 0 | 0
20 | HOAG 10202 | 25-209 | SPOT VITAL SIGN MONITOR | 13650 | >10 | 85% | Uptime % Apr: 99.99%; Uptime % May: 99.98% | 30 |  |  | N | N | 0 | 0 | 0
21 | HOAG 10110 | 11-165 | ADEC CHAIR BASIC 2140 AND STOOL | 49950 | >10 | 85% | Uptime % Apr: 99.9% | 100 |  |  | N | N | 0 | 0 | 0
22 | HOAG 10409 | 13-372 | RETINOSCOPE | 950 | 5-10 YRS | 92% | Uptime % Apr: 99.95% | 10 |  |  | N | N | 0 | 0 | 0
23 | HOAG 10236 | 12-610 | FETAL HEART MONITOR (DAPTONE) | 2400 | 5-10 YRS | 92% | Uptime % Apr: 99.99% | 10 |  |  | N | N | 0 | 0 | 0
24 | HOAG 10309 | 11-407 | ECG (EKG) / EPG PROJECT - ECG MACHINE | 5950 | >10 | 85% | Uptime % May: 99.96% | 20 |  |  | N | N | 0 | 0 | 0
25 | HOAG 10019 | 25-209 | PATIENT MONITORING EQUIP (MINDRAY) | 5795 | >10 | 85% | Uptime % May: 99.99% | 20 |  |  | N | N | 0 | 0 | 0
26 | HOAG 10063 | 25-209 | SPOT VITAL SIGN / TRANSPORT VITAL SIGN MONITOR | 8000 | >10 | 85% | Uptime % May: 99.95% | 20 |  |  | N | N | 0 | 0 | 0
27 | HOAG 10369 | 25-209 | PATIENT MONITORING EQUIPMENT | 8500 | 5-10 YRS | 92% | Uptime % May: 99.99% | 20 |  |  | N | N | 0 | 0 | 0
28 | HOAG 10289 | 17-148 | PATIENT MONITORING EQUIPMENT | 3000 | >10 | 88% | Uptime % May: 99.99% | 10 |  |  | N | N | 0 | 0 | 0
TOTAL DEDUCTION FOR IND. B.4: | [25 empty] | 0
24 | 55 | 1320 | 8760 | 0.84932 |  |  | Uptime % Sep: 90; Uptime % Oct: 8760; Uptime % Nov: 8670

## B5-T&C
MINISTRY OF HEALTH
Deduction Mapping Report Biomedical Engineering Maintenance Services (BEMS)
INDICATOR B.5 TESTING & COMMISSIONING AND WARRANTY MANAGEMENT
| HOSPITAL NAME : | HOSPITAL ORANG ASLI GOMBAK | [5 empty] | MONTH/YEAR : |  | 2026-05-01
| STATE : | WILAYAH PERSEKUTUAN KUALA LUMPUR | [5 empty] | BEMS MONTHLY FEE (RM) @ 97.5 % : |  | 17318.1547
| COMPANY : | AMBANG WIRA SDN BHD | [5 empty] | FMS MONTHLY FEE (RM) @ 2.5% : |  | 444.0553
Notes:
1) Deduction Ringgit Equivalent is based on Appendix 11 of the Agreement, Section B3 Key Deduction Indicators for BEMS
2) Deduction Value is Deduction Ringgit Equivalent (per event)
3) Display only asset affected by deduction
4) *May not be applicable for unregistered asset
5) **Only applicable for under warranty management deduction
No | Work Order/ NCR | Work Order/ NCR Date | Work Order/ NCR Reference No. | Work Order/ NCR Description | User Location | T&C Date | Completion Date | Asset Number* | Asset Description | Warranty Expiry Date** | Asset Purchase Price (RM) | Deduction Ringgit Equivalent (RM) | Validate Status (Y/N) | Deduction Value (RM) | Remarks
Work Order Close
1 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | - | N | - | NA
Work Order Open
1 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | - | N | - | NA
TOTAL DEDUCTION FOR IND. B.5: | [13 empty] | 0

## Sheet1
Equipments | Group | Operating Hours | Operating Hours/WK (Days) | Uptime Equipment | Uptime Equipment | Uptime Equipment
|  |  |  | < 5YRS | 5-10 YRS | >10 YRS
Therapeutics
Alternating Pressure Mattress | P | 12 | 6 | 96% | 92% | 85%
Aneasthesia Units and Vaporizers | C | 24 | 7 | 99% | 95% | 88%
Aneasthesia Ventilators | C | 24 | 7 | 99% | 95% | 88%
Aspirators | C | 24 | 7 | 99% | 95% | 88%
Defibrillators | C | 24 | 7 | 99% | 95% | 88%
Dental Equipment | P | 24 | 7 | 96% | 92% | 85%
Diathermy Units | C | 24 | 7 | 99% | 95% | 88%
Electrosurgical Units | C | 24 | 7 | 99% | 95% | 88%
Heamodialysis Units | C | 24 | 7 | 99% | 95% | 88%
Humidifiers | C | 24 | 7 | 99% | 95% | 88%
Hypo/hyperthermia Units | C | 24 | 7 | 99% | 95% | 88%
Incubators | C | 24 | 7 | 99% | 95% | 88%
Infusion controllers/pumps | C | 24 | 7 | 99% | 95% | 88%
Lasers | C | 24 | 7 | 99% | 95% | 88%
Lithotripters | C | 12 | 6 | 99% | 95% | 88%
Pacemakers | C | 24 | 7 | 99% | 95% | 88%
Peritoneal Dialysis Unit | C | 24 | 7 | 99% | 95% | 88%
Phototheraphy Units | C | 24 | 7 | 99% | 95% | 88%
Radiant Warmers | P | 24 | 7 | 96% | 92% | 85%
Radiographic Dye Injection | C | 24 | 7 | 99% | 95% | 88%
Radiotheraphy Equipments | P | 12 | 6 | 96% | 92% | 85%
Resuscitators | C | 24 | 7 | 99% | 95% | 88%
Stimulators | P | 12 | 6 | 96% | 92% | 85%
Surgical Drills and Saws | C | 24 | 7 | 99% | 95% | 88%
Surgical Tables | C | 24 | 7 | 99% | 95% | 88%
Traction Units | P | 12 | 6 | 96% | 92% | 85%
Ultrasonic Nebulizers | P | 12 | 7 | 96% | 92% | 85%
Ultrasonic Therapy | P | 12 | 6 | 96% | 92% | 85%
Ventilators | C | 24 | 7 | 99% | 95% | 88%
Ambulatory ECG Recorder and Scanners | P | 24 | 7 | 96% | 92% | 85%
Apnea Monitors | C | 24 | 7 | 99% | 95% | 88%
Blood Pressure Units (Non-invasive Inversive Electronic Units) | C | 12 | 6 | 96% | 92% | 88%
Capnographs | C | 24 | 7 | 99% | 95% | 88%
Cardiac Output Units | P | 24 | 7 | 96% | 92% | 85%
Diagnostic Radiologic Imaging Units | C | 24 | 7 | 99% | 95% | 88%
Electrocardiographs | P | 12 | 6 | 96% | 92% | 85%
Electroencephalographs | P | 12 | 6 | 96% | 92% | 85%
Electromyographs | P | 12 | 6 | 96% | 92% | 85%
Electronic Thermometers | P | 12 | 6 | 96% | 92% | 85%
Endoscopes / Bronschopes (Fiber Optic) | C | 24 | 7 | 96% | 92% | 88%
Evoked Potential Units | P | 12 | 6 | 96% | 92% | 85%
Fetal Monitors | C | 24 | 7 | 99% | 95% | 88%
Nuclear Medicine Equipments | P | 24 | 7 | 96% | 92% | 85%
Oximeters | C | 24 | 7 | 99% | 95% | 88%
Oxygen Monitors and Analysers | C | 24 | 7 | 99% | 95% | 88%
Phonocardiographs | P | 12 | 6 | 96% | 92% | 85%
Physiologic Monitoring Systems and Monitors | C | 24 | 7 | 99% | 95% | 88%
Pulmonary Function Analysers | P | 12 | 6 | 99% | 92% | 85%
Surgical Lights | C | 12 | 6 | 99% | 95% | 88%
Surgical Microscopes | C | 24 | 7 | 99% | 95% | 88%
Transcutaneous O2 and CO2 monitors | C | 24 | 7 | 99% | 95% | 88%
Treadmills | P | 12 | 6 | 96% | 92% | 85%
Ultrasound Imaging Units / Systems | P | 12 | 6 | 96% | 92% | 85%
Vector Cardiographs | P | 12 | 6 | 96% | 92% | 85%
Analysers
· Amino Acid | P | 12 | 6 | 96% | 92% | 85%
· Bilirubinometers | P | 12 | 6 | 96% | 92% | 85%
· Blood Gas | P | 12 | 6 | 96% | 92% | 85%
· Calcium | P | 12 | 6 | 96% | 92% | 85%
· Clinical Chemistry | P | 12 | 6 | 96% | 92% | 85%
· Coagulation | P | 12 | 6 | 96% | 92% | 85%
· Counter, Gamma | P | 12 | 6 | 96% | 92% | 85%
· Electrolyte | P | 12 | 6 | 96% | 92% | 85%
· Immunoassay | P | 12 | 6 | 96% | 92% | 85%
· Glucose | P | 12 | 6 | 96% | 92% | 85%
· Hematology | P | 12 | 6 | 96% | 92% | 85%
· Platelet Aggregation | P | 12 | 6 | 96% | 92% | 85%
Atomic Absorption Units | P | 12 | 6 | 96% | 92% | 85%
Automatic Blood Grouping Systems | P | 12 | 6 | 96% | 92% | 85%
Automatic Microbiological Systems | P | 12 | 6 | 96% | 92% | 85%
Blood Bank centrifuges | P | 12 | 6 | 96% | 92% | 85%
Centrifuges | P | 12 | 6 | 96% | 92% | 85%
Chromatographs Gas/Liquid | P | 12 | 6 | 96% | 92% | 85%
Cytometers | P | 12 | 6 | 96% | 92% | 85%
Electrophoresis Equipment | P | 12 | 6 | 96% | 92% | 85%
Flame Photometers | P | 12 | 6 | 96% | 92% | 85%
Microscopes | P | 12 | 6 | 96% | 92% | 85%
pH Meters | P | 12 | 6 | 96% | 92% | 85%
Spectrophotometers | P | 12 | 6 | 96% | 92% | 85%

## KDI
Indicator No. | Indicator | Frequency of Deduction | Biomedical Equipment Purchase Cost (RM) |  | Deduction Range Value (RM) |  | Maximum Demerit Point per Work Order
|  |  | From | To
B.1 | The Company shall respond to the request within the specified time frame | Per event | 0.00 5,000.00 10,000.00 15,000.00 20,000.00 25,000.00 30,000.00 35,000.00 40,000.00 > 45,000.00 | 4,999.99 9,999.99 14,999.99 19,999.99 24,999.99 29,999.99 34,999.99 39,999.99 44,999.99 | 10.00 20.00 30.00 40.00 50.00 60.00 70.00 80.00 90.00 100.00 |  | 1
B.2 | The Company shall carry out repairs to restore faulty medical equipment to a specified condition when it fails to function as per manufacturer’s specification within seven (7) working days. | Per day | 0.00 5,000.00 10,000.00 15,000.00 20,000.00 25,000.00 30,000.00 35,000.00 40,000.00 > 45,000.00 | 4,999.99 9,999.99 14,999.99 19,999.99 24,999.99 29,999.99 34,999.99 39,999.99 44,999.99 | 10.00 20.00 30.00 40.00 50.00 60.00 70.00 80.00 90.00 100.00 |  | Number of working days for the month
B.3 | The Company shall carry out scheduled maintenance as per schedule in accordance to the procedure as specified in Agreement. Scheduled maintenance shall include complete safety and performance tests. | Per event | 0.00 5,000.00 10,000.00 15,000.00 20,000.00 25,000.00 30,000.00 35,000.00 40,000.00 45,000.00 | 4,999.99 9,999.99 14,999.99 19,999.99 24,999.99 29,999.99 34,999.99 39,999.99 44,999.99 49,999.99 | 10.00 20.00 30.00 40.00 50.00 60.00 70.00 80.00 90.00 100.00 |  | 1
For every RM 5,000 increment of equipment cost, an increase of RM10 will be applied subsequently
B.4 | The Company shall ensure biomedical equipment meets the uptime target as set in the Agreement. | Per event |  |  | 1st level of Uptime Target (TRKPI) | 2nd level of Uptime Target (80%) | 2
|  |  | 0.00 5,000.00 10,000.00 15,000.00 20,000.00 25,000.00 30,000.00 35,000.00 40,000.00 45,000.00 | 4,999.99 9,999.99 14,999.99 19,999.99 24,999.99 29,999.99 34,999.99 39,999.99 44,999.99 49,999.99 | 10.00 20.00 30.00 40.00 50.00 60.00 70.00 80.00 90.00 100.00 | 20.00 40.00 60.00 80.00 100.00 120.00 140.00 160.00 180.00 200.00
For every RM 5,000 increment of equipment cost, an increase of RM10 will be applied subsequently for 1st Uptime Target and an increase of RM20 for 2nd Uptime Target will be applied subsequently
B.5 | The Company shall attend and witness T&C and manage biomedical equipment within the warranty period. | Per event | 0.00 5,000.00 10,000.00 15,000.00 20,000.00 25,000.00 30,000.00 35,000.00 40,000.00 45,000.00 | 4,999.99 9,999.99 14,999.99 19,999.99 24,999.99 29,999.99 34,999.99 39,999.99 44,999.99 49,999.99 | 10.00 20.00 30.00 40.00 50.00 60.00 70.00 80.00 90.00 100.00 |  | 1
For every RM 5,000 increment of equipment cost, an increase of RM10 will be applied subsequently

## Public Holiday
­
Tarikh | Hari | Cuti | Negeri Berkenaan
2026-01-01 | Khamis | Tahun Baru | Selangor
2026-02-17 | Selasa | Tahun Baru Cina | Selangor
2026-02-18 | Rabu | Tahun Baru Cina Hari Kedua | Selangor
2026-02-02 | Ahad | Hari Thaipusam | Selangor
2026-03-07 | Sabtu | Hari Nuzul Al-Quran | Selangor
2026-03-20 | Jumaat | Cuti Hari Raya | Selangor
2026-03-21 | Sabtu | Hari Raya Aidilfitri | Selangor
2026-03-22 | Ahad | Hari Raya Aidilfitri Hari Kedua | Selangor
2026-03-23 | Isnin | Cuti Hari Raya | Selangor
2026-05-01 | Khamis | Hari Pekerja | Selangor
2026-05-27 | Rabu | Hari Raya Haji | Selangor
2026-05-31 | Ahad | Hari Wesak | Selangor
2026-06-01 | Isnin | Hari Keputeraan YDP Agong | Selangor
2026-06-02 | Selasa | Cuti Hari Wesak | Selangor
2026-06-01 | Isnin | Hari Keputeraan YDP Agong | Selangor
2026-06-02 | Selasa | Cuti Hari Wesak | Selangor
2026-06-17 | Rabu | Awal Muharram | Selangor
2026-08-25 | Selasa | Maulidur Rasul | Selangor
2026-08-31 | Isnin | Hari Kebangsaan | Selangor
2026-09-16 | Rabu | Hari Malaysia | Selangor
2026-11-08 | Ahad | Hari Deepavali | Selangor
2026-11-09 | Isnin | Cuti Hari Deepavali | Selangor
2026-12-11 | Jumaat | Hari Keputeraan Sultan Selangor | Selangor
2026-12-25 | Jumaat | Hari Krismas | Selangor
* Subject to change (Tertakluk kepada perubahan), TBC - To be confirmed

## Sheet2 (hidden)
No | Asset No | Asset Description | Asset Purchase Price (RM) | Work Order No. |  | User Location | Work Order/ Complaint Details | Response Category (Emergency/ Normal) | Total Response Time (minutes) | Work Request Date | Work Completion Date | Deduction Ringgit Equivalent (RM) | Indicator B.1 |  |  | Indicator B.2 | [5 empty] | Contract
[5 empty] | wo | [7 empty] | Validate Status (Y/N) | Deduction Value (RM) | Remarks/ Exemption | Validate Status (Y/N) | Total Repair Time (day) | Last date of 7 working days (dd/mm/yy) | Demerit Point | Deduction Value (RM) | Remarks/ Exemption |  | Total Repair Time (day) | Hours | Total Hours | Total Hours / Year | Balance Uptime | Uptime Percentage | Uptime TRKPI | Uptime Deduction Yes / No
23 | IKN400966 | LINEAR ACCELERATOR | 12049370 | 1811 |  | RT159 | Lazer Double | Normal | 0:15 | 2019-11-12 | 2019-11-26 | 100 | N | 0 | Response On_Time | N | 15 | 2019-11-20 | 6 | 600 |  | Strands | 15 | 24 | 360 | 8072 | 7712 | 95.54% | 92% | No
1 | IKN400292 | Infusion Pumps, General-Purpose | 10825 | 1655 |  | WAD 8A | Infucient Pump-Mesin Problem | Normal | 0:37 | 2019-11-05 | 2019-11-30 | 30 | N | 0 | Response On_Time | Y | 26 | 2019-11-13 | 17 | 510 |  | Strands | 26 | 24 | 624 | 8072 | 7448 | 92.27% | 95% | Yes
2 | IKN400978 | Scanning Systems, Computed Tomography | 3557530 | 1172 |  | RT | sendingg image from vsim to mosaiq | Normal | 0:30 | 2019-11-01 | 2019-11-13 | 100 | N | 0 | Response On_Time | N | 13 | 2019-10-25 | 13 | 1300 |  | Strands | 13 | 24 | 312 | 8072 | 7760 | 96.13% | 95% | No
3 | IKN400973 | Scanning Systems, Computed Tomography | 3557530 | 1173 |  | RT | sendingg image from vsim to mosaiq | Normal | 0:20 | 2019-11-01 | 2019-11-13 | 100 | N | 0 | Response On_Time | N | 13 | 2019-10-25 | 13 | 1300 |  | Strands | 13 | 24 | 312 | 8072 | 7760 | 96.13% | 95% | No
4 | IKN400151 | Mattress Systems, Alternating-Pressure | 28630 | 1449 |  | WAD 7A | Motor ripple mattress rosak | Normal | 0:40 | 2019-11-01 | 2019-11-06 | 60 | N | 0 | Response On_Time | N | 6 | 2019-11-05 | 6 | 360 |  | Strands | 6 | 24 | 144 | 8072 | 7928 | 98.22% | 92% | No
5 | IKN400967 | Linear Accelerators | 12049370 | 1471 |  | RT | Table verticle movement problem | Normal | 1:35 | 2019-11-01 | 2019-11-04 | 100 | N | 0 | Response On_Time | N | 4 | 2019-11-06 | 4 | 400 |  | Strands | 4 | 24 | 96 | 8072 | 7976 | 98.81% | 92% | No
6 | NB 355296 | Mattress Systems, Alternating-Pressure | 28630 | 1476 |  | WAD 7B | Ripple Mattress tidak boleh pump | Normal | 0:34 | 2019-11-01 | 2019-11-06 | 60 | N | 0 | Response On_Time | N | 6 | 2019-11-06 | 6 | 360 |  | #N/A | 6 | 24 | 144 | 8072 | 7928 | 98.22% | 92% | No
7 | IKN400350 | Sphygmomanometers, Electronic, Automatic | 4510 | 1496 |  | WAD 7B | BP Set warna putih rosak. | Normal | 0:15 | 2019-11-01 | 2019-11-06 | 10 | N | 0 | Response On_Time | N | 6 | 2019-11-06 | 6 | 60 |  | Strands | 6 | 24 | 144 | 8072 | 7928 | 98.22% | 92% | No
8 | IKN400775 | Scanning Systems, Ultrasonic, General-Purpose | 1044560 | 1535 |  | ID | Ultrasound tidak keluar gambar | Normal | 0:25 | 2019-11-01 | 2019-11-04 | 100 | N | 0 | Response On_Time | N | 4 | 2019-11-07 | 4 | 400 |  | Strands | 4 | 24 | 96 | 8072 | 7976 | 98.81% | 95% | No
9 | IKN400567 | Cameras, Gamma | 3191750 | 1550 |  | NM | MEGP detector one collision | Normal | 0:05 | 2019-11-01 | 2019-11-04 | 100 | N | 0 | Response On_Time | N | 4 | 2019-11-08 | 4 | 400 |  | Strands | 4 | 24 | 96 | 8072 | 7976 | 98.81% | 92% | No
10 | IKN400786 | Radiographic/Fluoroscopic Systems, Angiographic/Intervention | 3191750 | 1565 |  | ID | Request for user training at Radiology Department (Angio) on 23/10/2019. | Normal | 0:00 | 2019-11-01 | 2019-11-04 | 100 | N | 0 | Response On_Time | N | 4 | 2019-11-08 | 4 | 400 |  | Strands | 4 | 24 | 96 | 8072 | 7976 | 98.81% | 92% | No
6090
Formulas:
Last date of 7 working days (dd/mm/yy) = WORKDAY([Work Request Date],6,tarikhcuti)